Table 1. Method-comparison metrics for the four static scenarios (n = 150 paired measurements each).

| Scenario | Mean bias (mmHg) | 95% LoA (mmHg) | PE (%) | BA reg. slope (95% CI) | Deming slope (95% CI) | P-B slope (95% CI) | CCC | C_b | v |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| S1 offset only | +11.9 | 5.6 to 18.2 | 4.6 | 0.010 (-0.007, 0.028) | 1.010 (0.993, 1.028) | 1.010 (0.990, 1.029) | 0.921 | 0.926 | 1.010 |
| S2 zeroed (ideal) | +0.1 | -5.4 to 5.5 | 4.2 | -0.010 (-0.025, 0.006) | 0.990 (0.975, 1.006) | 0.985 (0.971, 1.001) | 0.995 | 1.000 | 0.990 |
| S3 gain error | +12.7 | 4.7 to 20.7 | 5.9 | 0.094 (0.080, 0.109) | 1.099 (1.082, 1.117) | 1.100 (1.081, 1.117) | 0.915 | 0.919 | 1.099 |
| S4 gain masked | +0.3 | -8.4 to 8.9 | 6.7 | 0.103 (0.087, 0.119) | 1.109 (1.091, 1.126) | 1.111 (1.092, 1.130) | 0.990 | 0.995 | 1.108 |

BA = Bland-Altman; CI = confidence interval; LoA = limits of agreement; PE = percentage error; P-B = Passing-Bablok; CCC = concordance correlation coefficient; C_b = bias-correction factor; v = scale shift.

Table 2. Dynamic-response metrics for pulse pressure (PP) and systolic pressure (SBP).

| System (f_n, zeta) | PP ratio | PP mean bias (mmHg) | CCC (PP) | v (PP) | SBP mean bias (mmHg) |
| --- | --- | --- | --- | --- | --- |
| Optimal (25.0 Hz, 0.65) | 1.01 | +0.3 | 0.973 | 1.030 | +0.2 |
| Underdamped (10.0 Hz, 0.15) | 1.24 | +10.7 | 0.588 | 1.277 | +7.6 |
| Overdamped (8.0 Hz, 0.80) | 0.92 | -3.4 | 0.891 | 0.989 | -3.4 |

PP ratio = mean(measured PP / true PP); f_n = natural frequency; zeta = damping coefficient.

Table 3. Real-waveform validation metrics for the four static scenarios (n = 1,874 paired beats per scenario from the VitalDB Open Dataset, 29 cases).

| Scenario | n | Mean bias (mmHg) | 95% LoA (mmHg) | PE (%) | BA reg. slope (95% CI) | Deming slope (95% CI) | P-B slope (95% CI) | CCC | C_b | v |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| R1 offset only | 1874 | +12.0 | 6.1 to 17.9 | 4.5 | 0.006 (0.001, 0.012) | 1.006 (1.001, 1.012) | 1.006 (1.000, 1.011) | 0.894 | 0.900 | 1.006 |
| R2 zeroed (ideal) | 1874 | -0.0 | -5.9 to 5.9 | 4.7 | 0.006 (0.001, 0.012) | 1.006 (1.001, 1.012) | 1.006 (1.000, 1.011) | 0.993 | 1.000 | 1.006 |
| R3 gain error | 1874 | +12.6 | 4.9 to 20.3 | 5.8 | 0.101 (0.096, 0.106) | 1.106 (1.101, 1.112) | 1.105 (1.100, 1.111) | 0.890 | 0.896 | 1.106 |
| R4 gain masked | 1874 | -0.0 | -7.7 to 7.7 | 6.1 | 0.101 (0.096, 0.106) | 1.106 (1.101, 1.112) | 1.105 (1.100, 1.111) | 0.989 | 0.995 | 1.106 |

BA = Bland-Altman; CI = confidence interval; LoA = limits of agreement; PE = percentage error; P-B = Passing-Bablok; CCC = concordance correlation coefficient; C_b = bias-correction factor; v = scale shift.